## RFQ
Item requested |  |  |  |  | Item offered | [7 empty] | Logistical information |  |  |  | Manufacturer |  |  |  | Stock and Lead time information |  |  | Maintenance and Service arrangements |  | Remarks
SeqNo | Item Description | UOM | Quantity | Requested delivery date | Description | UOM | Quantity | EXW Unit Price ($) | Total Price ($) | Installation and Training fee ($);that is if not included in unit price;otherwise fill zero | Warranty:Y/N (also indicate period in years) | Attach/Link to product information sheet | EXW Place | Instrument Manual Label Language (s) | Shipping Weight | Shipping Volume | Name | Address plant | Manufactured in a ISO compliant site (Y/N)? | Item is SRA approved (CE marked, FDA etc)? | Quantity on stock | Lead time stock | Lead time fresh production | Minimum recommended Service and maintenance (once per year, twice per year etc) | **Estimated annual maintenance and service costs ($) including spare parts, calibration, system updates etc
1 | High Pressure Liquid Chromatography (HPLC)-with UV/VIS System | each | 4 | 2018-04-30
2 | UV/VIS Spectrophotometer | each | 3 | 2018-04-30
3 | FT-IR Spectrophotometer | each | 1 | 2018-04-30
4 | Automated Volumetric Karl Fisher Titrator (potentiometer) | each | 3 | 2018-04-30
5 | Micro-Analytical Balance with printer | each | 1 | 2018-04-30
6 | pH Meter | each | 2 | 2018-04-30
7 | Tablet Friability tester | each | 2 | 2018-04-30
8 | Weighing Booth | each | 1 | 2018-04-30
9 | Analytical Balance with printer | each | 5 | 2018-04-30
10 | Ultrasonic Bath (sonnicator) | each | 2 | 2018-04-30
11 | Water De-ionizer | each | 1 | 2018-04-30
12 | Dissolution Apparatus | each | 4 | 2018-04-30
13 | Gas Chromatography-FID System | each | 1 | 2018-04-30
14 | Water Bath | each | 1 | 2018-04-30
15 | Polarimeter | each | 1 | 2018-04-30
16 | Disintergration Apparatus | each | 3 | 2018-04-30
17 | Electric Oven | each | 1 | 2018-04-30
18 | Heavy Duty UPS for Laboratory Equipment | each | 1 | 2018-04-30
* PFSCM's Terms & Conditions are applicable to all of PFSCM's procurements. Quoting in response to this RFQ automatically means that PFSCM's Terms & Conditions are accepted by the vendor.
* Installation training includes Installation qualification, Operation qualification and Performance qualification (IP, OP &PQ )
** Exclude travel costs to Malawi in the estimated annual maintenance and service fee

## Annex B - Compliance sheet
Requested item description |  | Offered item description | Comply:Y/N
1. High Pressure Liquid Chromatography (HPLC)-with UV/VIS System
PROPERTY | DESCRIPTION
System and Performance
1.1 General | A HPLC system with quaternary reciprocating pumping system capable of solvent delivery in isocratic and gradient modes. The system should be equipped with a UV/VIS detector,Autosampler , Column oven and degasser with microcomputer electronics and PC control
1.2 Pump flow-rate | Minimum 0.0l ml/ minute, Maximum 10ml/minute, Precision 0.1%RSD, Linearity 0.9999
1.3 Pump Pressure | 0 to 600Bars
1.4 Column oven | Temperature up to 60 °C
1.5 Detector | UV/VIS with a wavelength range of 190nm to 700nm, wavelength accuracy: ±0.1, Wavelength repeatability: ±0.05
1.6 Autosampler | Injection volume: 5 - 20µl , Minimum number of vials per tray: 100
1.8 Degasser | In-line degasser
2.0 Operating condit ions | 15°c to 35°C, <75% humidity
3.0 Power requirements | 220 to 250AC, 50/ 60Hz
4.0 Data System and software
4.1 General | A software platform that incorporates all functions required for HPLC-UV/VIS analysis with instrument control, data manipulation and analysis with flexible report utilities
4.2 User interface | User login function(password-protected). Supervisor controlled access to methods and routines,menu, tool bar and tool box functions
4.3 Reports | Quick print facility for chromatograms and results window
4.4 Other | Supplier to provide installation training including Installation qualification, Operation qualification and Performance qualification (IP, OP &PQ ). In addit ion supplier to provide Autosampler vials (5000), Vacuum filter bottles(2) and 2 Columns(Type of column :C18 , 25 cm x 4.6 mm, 5µm)
2. UV/VIS Spectrophotometer
PROPERTY | DESCRIPTION
Optical System & Performance
1.1 General | Double beam UV/VIS SPectrophotometer with micro-computer electronics and PC control
1.2 Acqu isition speed | ≤ 0.1S
1.3 Wave length r ange | 190 -1100nm
1.4 Stray light | < 0.01%
1.5 Wavelength accuracy | ±0.1
1.6 Wavelength repeatability | ±0.05
1.7 Photometric accuracy at 1A | ±0.001
1.8 Photometric accuracy at 2A | ±0.005
Operating conditions | 15°C to 35°C, <75% humidity
Power requirements | 220 to 250AC, S0/ 60Hz
Data System and Software
2.0 General | A software platform that incorporates all functions required for UV/VIS analysis including wavelength scanning and fixed wavelength operations
2.3 Reports | Quick print facility for graphs, spectra and results window
Other | Supplier to provide installation training including Installation qualification, Operation qualification and Performance qualification (IP, OP &PQ )
3. FT-IR Spectrophotometer
PROPERTY | DESCRIPTION
Optical System & Performance | Desiccated and sealed optical unit with vibration isolated base- plate capable of handling samples on KBr discs as well as samples on a diamond plate (ATR mode)
1.1 General
1.2 Spectral Resolution | 1.0 cm¯1
1.3 Wave number range | 7500-400cm¯1
1.4 Signal to Noise ratio | 30,000:1
1.5 Detector | Electrically, temperature-stabilized fast recovery deuterated triglycine Sulfate(FR-DTGS) or Lithium Tantalate(LiTa03)
1.6 Beam splitters | Optimized multilayer Potassium Bromide
1.7 Wavenumber accuracy | 0.5 cm¯1
1.8 Wavenumber repeatability | ±0.0009 cm¯1
1.9 Light Source | User replaceable long-life source
2.0 Interferometer | Improved Michelson Interferometer, self-compensating for dynamic alignment changes due to tilt and shear, incorporating high reflectivity, and first-surface aluminium -coated optics
2.1 Support Accessories | One hand-driven pressfor the producti on of KBr pellets (discs)
Data System and Software
2.1 General | A software platform that incorporates all functions required for infrared analyses: instrument control(including self-diagnostic f unction), data manipulation and analysis with flexible report utilities
2.2 User interface | User login function (password-protected). Supervisor controlled access to methods and routines, menu, toolbarandtool box functions
2.3 Reports | Quick print facility for graphs, spectra and result s window
2.4 Material testing | Spectral comparison algorithms for searching against commercially available or customer developed libraries including capacity to validate instrument performance
Operating Environment
2.5 Temperature | 10 to 35°C
2.6 Humidity | 20 to 70% RH
Other | Supplier to provide installation training including Installation qualification, Operation qualification and Performance qualification (IP, OP &PQ )
4. Automated Volumetric Karl Fisher Titrator
PROPERTY | DESCRIPTION
General | Automatic volumetric Karl Fisher Titrator capable of determining water content from 100ppm to 100% in a wide range of solid, liquid and gaseous samples. The instrument package should include a magnetic stirrer with a stirring rod, polarizable electrode(double pin Platinum) with cable, 5ml burette, titration vessel, solvent manager pump for emptying/filling vessels, method templates, one desiccant bottle and other necessary accessories required to make it fully functional
Operating conditions | 10°c to 45°c, ≥85% humidity
Power requirements | 220 to 250AC, 50/60Hz
Measurement range | -2000 to 2000 mV
Resolution | 0.1mV
Measuring Accuracy | ±0.2mV
Polarization current range | -130 to +110µA
Polarization voltage range | -1200 to+ 1200mV
Measurement range(Water content) | 100ppm to 100%
Dosing Accuracy | ±0.3%
Methods and Series | Minimum number of pre-programmed methods for most important applications ≥2
Control and Evaluation | End point titration with dynamic and continuous titration
Equivalence point titration with dynamic or incremental titrant addition
Predefined formulas for the most common calculations
Stirrer | Internal magnetic stirrer or propeller stirrer
Minimum Burette Volume | 5ml
Data System and soft ware | Software providing links to PC, Keyboard, USB printer, barcode reader,Balance and USB flash drives
Other | Supplier to provide installation training including Installation qualification, Operation qualification and Performance qualification (IP, OP &PQ )
5. Micro-Analytical Balance
PROPERTY | DESCRIPTION
General | A weighing instrument featuring an automatic time and temperature control overload protection safeguards, a visual display with in-built menu for simple cover, programmable keys for short-cut access to preferred applications and allowing convenient transfer of weighing results. Additionally, the balance should be of good quality chemical-resistant material that will protect sensitive sensors from environmental effects/impacts.
Operating conditions | 0°C to 40°C, <75% humidity
Power requirements | 220 to 250AC, 50/60Hz
Maximum Capacity | 20 to 30g
Readability | 0.001mg
Minimum weight according to USP | ≤ 1.5mg
Repeatability(for 20g) | ≤ 0.0025mg (as standard deviation)
Repeatability for 5% load | ≤ 0.0015mg (as standard deviation)
Tare range | 0 to 30g
Sensitivity temperature drift(0 to 30°c) | 0.0001% per degree Celcius
Sensitivity stability per annum | 0.0001% per annum
Settling time | ≤ 5 Seconds
Eccentricity (for 10g} | ≤ 0.02mg
Calibration type | Internal
Linearity deviation | ≤ 0.006mg
Language | English
Weighing units | mg
Data System and software | Data printing capability with GLP and GMP function
Other | Supplier to provide installation training including Installation qualification, Operation qualification and Performance qualification (IP, OP &PQ )
6. pH Meter
PROPERTY | DESCRIPTION
General | An instrument capable of measuring pH, potential and temperature. The apparatus should offer GLP compliant electrode test yielding specific information such as zero point or slope, response time as well as a statement that informs the user the conditions of the membrane and diaphragm.The instrument package should include a suitable electrode, magnetic stirrer, LCD and all accessories required for full functionality
Operatingconditions | 10°c to 40°c, ≥75% humidity
Power requirements | 220 to 250AC, 50/60Hz
Measurement range(pH) | 0 to 14
Measurement range (Potential) | ±2200mV
Measurement range(Temperature) | -20 to 250°C
Resolution | pH= 0.001, Potential = 0.1mV, Temperature = 0.1°c
System diagnosis | Internal
Calibration | pH calibration with up to six or more buffers offering calibration data/curves storage
LCD Display | Should display important data (slope, pH, temperature, date, time and calibration curves)
Magnetic stirrer | Internal magnetic stirrer controlled by the pH meter with speed control facility
Printer | Customized printer for printing information on the display including calibration curves, user identification details and field for signature on report. (GLP com pliant result printout)
Memory | Memory for 50 measured values and related data
Data System | Connection for PC Keyboard, printer, and barcode reader.
Other | Supplier to provide installation training including Installation qualification, Operation qualification and Performance qualification (IP, OP &PQ )
7. Tablet Friability
PROPERTY | DESCRIPTION
General | A double drum USP/EP/JP compliant Friability/ Abrasion tester with a programmable speed of between 20 &100 RPM and a USB interface allowing test parameters (speed, time) download to a PC or receipt of a hard copy print -out when connected to a printer
Operating conditions | 10°C to 40°C,≥75% humidity
Power requirements | 220 to 250AC, 50/60Hz
M easurement speed | 20 to 100RPM
Measurement speed accuracy | ±1RPM
No of drums | 2 ( which can be put into use simultaneously)
Sample discharge | Automatic after cycle completion
Timer | Integrated timer for precise setting of run time
Test run timer | Infinitely adjustable (≥5hrs, 59mins , 59secs)
Timer accuracy | ±0.1Secs
Revoluti ons counter | Infinitely adjustable (up to 9999)
LCD Display | Should display test parameter (RPM, Time)
Printer | Customized printer for printing information on the display
Other requirements | Provision of 10° tilting of drums asper USP requirement.Instrument should include 2 friability drums (Left & right) andone abrasion drum
Supplier to provide installation training including Installation qualification, Operation qualification and Performance qualification (IP, OP &PQ )
8. Weighing Booth
PROPERTY | DESCRIPTION
General | A weighing booth/ station designed for powder containment applications, delivering the proper airflow at the face of the enclosure to protect the operator from exposure to hazardous powders or active pharmaceutical compounds. Its enclosure should aim at reducing airflow turbulence and maintain balance stability to 4 decimal places without compromising containment. The air exhausted into the room should comply with class 5 EN ISO 14644
Operating conditions | 0°C to 40°C,˂75% humidity
Power requirements | 220 to 250AC, 50/60Hz
Internal Dimensions | Width: 1220 to 1250mm, Height: 800 to 900mm
Airflow meter | Permanent air face velocity monitoring system
Flow monitor | Permanent ventilation control device
Internal lighting | Compact tubular fluorescent lighting
Sampling port | For filter saturation detection
Programmable timer | Timer 40 to 60 hours setting by default
Chemical listing | Provide a list of chemicals which the station is capable of handling
Other requirements | Provide a metallic fixed work bench equipped with height adjusting jacks. Equipped with two height adjusting jacks.
Supplier to provide installation training including Installation qualification, Operation qualification and Performance qualification (IP, OP &PQ )
9. Analytical Balance with printer
PROPERTY | DESCRIPTION
General | A weighing instrument featuring an automatic time and temperature controlled internal calibration adjustment, overload protection safeguards, a visual display with in-built menu for simple operation, draft shield, protective in-use cover, programmable keys for short-cut access to preferred applicationsand direct connection to a pr int er or USB slot allowing convenient transfer of weighing result s. Additionally, the balance should be made of robust housing made of good quality chemical-resistant material that will protect sensitive sensors from environmental effects/impacts
Operating conditions | 0°C to 40°C,˂75% humidity
Power requirements | 220 to 250AC, 50/60Hz
M a ximum Capacit y | 200 to 250g
Readability | 0.01mg
M inimum weight according to USP | ≤20mg
Repeatability(for 100g or 200g) | ≤0.05mg
Tare range | 0 to 240g
Sensitivit y temperature drift (0 to 30°C ) | 0.0001% per degree Celcius
Sensitivity stability per annum | 0.0001% per annum
Settling time | ≤5 Seconds
Eccentricit y (for 100g) | ≤0.2mg
Calibrat ion t ype | Internal
Linearity deviation | ≤0.2mg
Language | English
Weighing units | g, mg
Data System and software | Data printing capability with GLP and GMP function
Other | Supplier to provide installation training including Installation qualification, Operation qualification and Performance qualification (IP, OP &PQ )
10. Ultrasonic Bath
PROPERTY | DESCRIPTION
General | A stain-less steel ultrasonic bath with electronic time and temperature control providing a time and or temperature pre-selection accessory/function.The instrument should have auto-degas/degas as well as sweep functions. It should be capable of ultrasound cleaning of laboratory glassware or metal, degassing HPLC solutions, mixing,dispersing, emulsifying and dissolving samples
Operating conditions | 0°C to 40°C,˂75% humidity
Power requirements | 220 to 250AC, 50/60Hz
Temperature | 30°C to 80°C (in not more than 5°celsiusincrement al steps)
Tank Volume | 10 to 15 Litres
Ultrasonic Frequency | ≥35KHz
Heating Power | ≥800W
Basket Load | ≥7Kgs
Safety | Auto-switch off function at a preset time for safety
LED Display | To display set cleaning period and remaining time as well as bath temperature
Basket | With 9x1mm mesh size and capable of holding four 1 litre flasks or more
Basket | With 9x1mm mesh size and capable of holding four 1 litre flasks or more
Other | Supplier to provide installation training including Installation qualification, Operation qualification and Performance qualification (IP, OP &PQ )
11. Water De-ionizer
PROPERTY | DESCRIPTION
General | An equipment combining water purification and monitoring technologies allowing the direct production of ultrapure water from portable water and quality monitoring of the same in a single device. The equipment should have the capacity to measure the total organic content and resistivity of its output.
Operating conditions | 0°C to 40°C, <75% humidity
Power requirements | 220 t o 250AC, S0/60Hz
Resistivity | ≥18.2MΩ.cm
Total Organic Content | ≤10ppb
Particulates (size :≥0.22µm) | <1particulate/ml
Bacteria | <0.01CFU/mL
Pyrogens | <0.001EU/mL
RNases | <1pg/mL
DNases | <5pg/mL
Flow rate | ≥2Ml/min
Production rate | ≥5litres/hour
Other | Supplier to provide installation training including Installation qualification, Operation qualification and Performance qualification (IP, OP &PQ )
12. Dissolution Tester
PROPERTY | DESCRIPTION
General | An 8-position, 8-Spindle configured dissolution tester complying with USP, BP andJP requirements. A dissolution apparatus supplied with a t ablet -droppingtray and capable of interface with other instruments such as UV/VIS or HPLC Systems
Bath Construction | Clear seamless plexiglass
Speed range | 25 to 250 ±0.1 RPM
Temperature range | Room temperature(Ambient) to 45° Celsius
Heater | ≥700W
Power requirements | 220 to 250AC, 50/ 60Hz
Shafts | Basket and Peddle interchangeable
Drive Positions | 8
Calibration | Self-calibrating speed and temperature control system
Interface with other Instruments | Automation ready for on-line analysis with UV/VIS, HPLC and fiber optic systems
Display | ±Easy to re ad LCD Display for speed, temperature and time
Other requirements | Provide 1000ml Glass Vessels(8), Shafts(S), Paddles(S), 40 mesh basket assembly(8), and a calibration ball(1)
Supplier to provide installation training including Installation qualification, Operation qualification and Performance qualification (IP, OP &PQ )
13. GC-FID System
PROPERTY | DESCRIPTION
System & Performance
1.1 General | A Gas Chromatograph with an FID detector, Autosampler, a programmable temperature vaporizer and an injector capable of handling Split/ Split less,direct and on-column injection modes with microcomputer electronics and PC control. The instrument should provide tool free inlet maintenance and column installation
1.2 Operating conditions | 15°C to 35°C, <75% humidity
1.3 Pow er requirements | 220 to 250AC, 50/ 60Hz
1.4 Column Oven | Temperature range: Maximum 450°C;Temperature Accuracy: Set value{K) ±1%; Temperature deviation: 2°C;Temperature variation coefficient: 0.01°C/°C; Temperature program steps:≥20; Cooling rate:300 to 50°C in less than 6 minutes; Acceptable columns: Capillary as well as Packed columns
1.5 Sample Injection unit | Temperature range: ≥400 °C; Sample injection modes: Split/ Split- less. On-Column, direct &programmable temperature vaporizer
1.6 Detector | Detector type: Flame Ionization Detector (FID); System: Dual flow rate differential system; Temperature range: ≥ 400°C; Limit of detection:≤3pg; Dynamic range:≥10⁷
1.7 Carrier Gas flow controller | Should provide constant flow rate control, linear speed control and pressure control; Maximum Pressure: ≥950KPa; Maximum Flow rate: 1,300Ml/minute
1.8 Data System and Software | General: A software platform that incorporates all functions required for GC-FID analysis with instrument control, data manipulation and analysis with flexible report utilities; User Interface: User login function (password-protected). Supervisor controlled access to methods and routines, menu , tool bar and tool box functions; Reports: Quick print facility for chromatograms and results window
Other | Supplier to provide installation training including Installation qualification, Operation qualification and Performance qualification (IP, OP &PQ}. In addition supplier to provide autosampler vials (5000 }, and 2 capillary Columns
14. Water Bath
PROPERTY | DESCRIPTION
General | A high performance general purpose water bath with capacity to maintain water temperatures from ambient to 100 degrees Celsius .The water bath should have over-temperatu re safety circuitry designed to prevent thermal runaway and auto-on and auto-off timers to allow for optimization of operation schedules
Chamber Capacity | 10L
Temperature range | Ambient to 100 degrees Celsius
Temperature stability | ±0.1degrees Celsius
Heater output | ≥800W
Power requirements | 220 to 250VAC, 50/ 60Hz
Operating conditions | 15°c to 35°c, <75%humidity
Other | Supplier to provide installation training including Installation qualification, Operation qualification and Performance qualification (IP, OP &PQ )
15: Polarimeter
PROPERTY | DESCRIPTION
General | A Polarimeter equipped with a sodium or Tungsten-Halogen Lamp capable of giving readings to the nearest 0.001 degrees, supplied with 2decimeter tubes
Operating conditions | 0°C to 40°C, <75% humidity
Power requirements | 220 to 250AC, 50/60Hz
Optical Wave length | 598nm
Sample Chamber | Capable of handling 2dm sample tubes
Sample tubes | Cylindrical glass cells
Measuring range | +90 degrees
Resolution | 0.01 degrees
Accuracy | ±0.01 degrees
Response speed | ±6 Degrees per second
Detector | PMT Tube
Readout modes | Optical rotation, Optical specific rotation, Concentration, Optical purity
Display | PC via USB ports (suitable software provided)
Calibration | In-built
Measurement temperature range | 0°C to 40°C (Accuracy ±0.1 Degree Celsius)
Other | Supplier to provide installation training including Installation qualification, Operation qualification and Performance qualification (IP, OP &PQ )
16. Disintergration Apparatus
PROPERTY | DESCRIPTION
General | A tablet disintegration apparatus compliant to USP/BP general chapters on disintegration testing
Operating conditions | 0°C to 40°C, ≥75% humidity
Power requirements | 220 to 250AC, 50/60Hz
Stroke Frequency | 30 ± 1 strokes/min
Stroke Height | 55 ± 2 strokes height
Number of Basket/Stations | 2 (each station driven independently
Temperature range | 30°C to 50°C
Integrated flow-through heating system accuracy: | ±0.2°C
LED Display | LED display/s to show remaining run-time (in minutes and seconds) and actual bath temperature
Other | Supplier to provide installation training including Installation qualification, Operation qualification and Performance qualification (IP, OP &PQ )
17. Electric oven
General: | An electric drying oven with a maximum heating temperature of 2500C
Power requirements: | 220 to 250AC, 50/60Hz
Operating conditions: | 100C to 400C, ≥75% humidity
Interior Volume: | 24Litres
Maximum Temperature: | 2500C
Temperature Variation: | ± 1°C
Readability: | 1°C
Heating ramp rate: | 250°C in 40 minutes
Display: | LED (displaying set temperature)
Alarm limit: | Settable
Other | Supplier to provide installation training including Installation qualification, Operation qualification and Performance qualification (IP, OP &PQ )
18. Heavy Duty UPS for Laboratory Equipment
PROPERTY | DESCRIPTION
Capacity | 2KVA, 2.1KW
Input | 230V AC, 50HZ
Current | 16 Amp
Output | 230V AC, 50HZ, 13, AMP
Other | Supplier to provide installation training including Installation qualification, Operation qualification and Performance qualification (IP, OP &PQ )